## Slide 1: Towards a verified cardiac pacemaker
Asankhaya Sharma Department of Computer Science National University of Singapore

## Slide 2: Introduction
Pacemaker
  Medical device which uses electrical impulses to fix abnormal heart
Pacemaker Formal Methods Challenge
  Based on a released informal specification of a previous generation of pacemaker by Boston Scientific
Related Work
  H Macedo, P Larsen, and J Fitzgerald, Incremental development of a distributed real-time model of a cardiac pacing system using vdm, FM 2008
  A Gomes and M Oliveira, Formal specification of a cardiac pacing system, FM 2009
  L. A. Tuan, M. C. Zheng, and Q. T. Tho, Modeling and Verification of Safety Critical Systems: A Case Study on Pacemaker, SSIRI, 2010

## Slide 3: Modeling with SPIN
Formal Specification in PROMELA
Verification of LTL Properties using SPIN

## Slide 4: Sequential Model
Update Timers
Heart
Sensor
Pace Generator
Global Variables

## Slide 5: Process - Update Timers
Clock Updates
  Increment Timer
  Reset Timer
Reset other Global variables
  AVD
  Pulses
  Senses

## Slide 6: Process - Heart
Different Behaviors
  Normal
    Wait NR, Pace A, Wait AVD, Pace V, Repeat
  Miss Ventricle Pace
    Wait NR, Pace A, Wait AVD, Skip, Repeat
  Dead
    Wait NR, Skip, Wait AVD, Skip, Repeat
  Non Deterministic
    Wait NR, May Pace A, Wait AVD, May Pace V, Repeat

## Slide 7: Process - Sensor
Record Paces from Heart and Pace Generator
Update Sense Variables
  Pace A -> Sense A
  Pace V -> Sense V
Capture Time of Senses for Refractory Period
  ARP
  VRP
  PVARP

## Slide 8: Process - Pace Generator
Modes Implemented
  VOO, AOO, DOO
  VVI, AAI, DDI
  VVT, AAT
  VDD, DDD

## Slide 9: VOO Mode
1
2
3
Min Time
Pace V
Reset Time

## Slide 10: VVI Mode
1
2
3
Min Time
Pace V
Reset Time
Sense V

## Slide 11: VVT Mode
1
2
3
Min Time
Pace V
Reset Time
Sense V
4
Pace V

## Slide 12: VDD Mode
1
2
3
Min Time
Pace V
Reset Time
4
Sense A
AVD Time
Sense V
5
Pace V

## Slide 13: LTL Properties
Deadlock
Pace Limit – LRLURLA and LRLURLV
AV Delay – AVD
Refractory Period – ARP, VRP and PVARP
Inhibiting – AAI and VVI
Triggering – AAT and VVT
Tracking – XDD

## Slide 14: Verification Results
LTL Property | VOO | AOO | DOO | VVI | AAI | DDI | VVT | AAT | VDD | DDD
Deadlock | X | X | X | X | X | X | X | X | X | X
Pace Limit | X | X | X | X | X | X | X | X | X | X
AV Delay |  |  | X |  |  | X |  |  | X | X
Refractory Period | X | X | X | X | X | X | X | X | X | X
Inhibiting |  |  |  | X | X | X |  |  |  | 
Triggering |  |  |  |  |  |  | X | X |  | 
Tracking |  |  |  |  |  |  |  |  | X | X

## Slide 15: Sequential Model (with Rate Control)
Update Timers
Heart
Sensor
Accelerometer
Global Variables
Pace Generator
Rate Controller

## Slide 16: Process - Accelerometer
Senses motion of Body
Readings correspond to Activity Threshold
  Very Low
  Low
  Mid
  High
  Very High

## Slide 17: Process - Rate Controller
Uses Activity Threshold to Change Rate of Pacing
  Response Factor
    Activity Threshold == Low -> RF = 1
    Activity Threshold == Med -> RF = 5
    Activity Threshold == High -> RF = 9
Calculate Rate of Pacing
  RF and Increment

## Slide 18: Hysteresis Mode
Valid for modes XXIX and XXDX
  Inhibiting or Tracking
Sense -> Wait (Regardless of Rate) -> Pace
  Can be simulated using RF

## Slide 19: Process - Pace Generator
Modes Implemented
  VOOR, AOOR, DOOR
  VVIR, AAIR, DDIR
  VDDR, DDDR
  VVI_H, AAI_H, DDI_H
  VDD_H, DDD_H

## Slide 20: VDDR Mode
1
2
3
Min Time + RF*Increment
Pace V
Reset Time
4
Sense A
AVD Time
Sense V
5
Pace V

## Slide 21: VDDRH Mode
1
2
3
Min Time + RF*Increment
Pace V
Reset Time
4
Sense A
AVD Time + RF*Increment
Sense V
5
Pace V

## Slide 22: LTL Properties
Rate Limit – LRLURLA_R and LRLURLV_R
Rate Control – LRLURLA_RC and LRLURLV_RC
Hysteresis Limit – AAI_H, VVI_H and XDD_H

## Slide 23: Verification Results
LTL Property | VOOR | AOOR | DOOR | VVIR | AAI R | DDI R | VDD R | DDD R | VV I H | A A I H | DDIH | VDDH | DDDH
Rate Limit | X | X | X | X | X | X | X | X |  |  |  |  | 
Rate Control | X | X | X | X | X | X | X | X |  |  |  |  | 
Hysteresis Limit |  |  |  |  |  |  |  |  | X | X | X | X | X

## Slide 24: Concurrent Model
Update Timers
Heart
Sensor
Pace Generator
Timing Variables
Pacing Variables
Sensing Variables
Guard Timing Variables
Guard Sensing Variables

## Slide 25: Verification Results
LTL Property | VOO | AOO | DOO | VVI | AAI | DDI | VVT | AAT | VDD | DDD
Deadlock | X | X | X | X | X | X | X | X | X | X
Pace Limit | X | X | X | X | X | X | X | X | X | X
AV Delay |  |  | X |  |  | X |  |  | X | X

## Slide 26: Distributed Model
Heart
Sensor
Pace Generator
Local Variables
Local Variables
Local Variables
Synchronization
Pulse
AVD
sense

## Slide 27: Verification Results
LTL Property | VOO | AOO | DOO | VVI | AAI | DDI | VVT | AAT | VDD | DDD
Deadlock | X | X | X | X | X | X | X | X | X | X
Pace Limit | X | X | X | X | X | X | X | X | X | X
AV Delay |  |  | X |  |  | X |  |  | X | X
Distributed AV Delay |  |  | X |  |  | X |  |  | X | X

## Slide 28: Conclusions and Future Work
Formal Specification of Cardiac Pace Maker in PROMELA
  Sequential
  Concurrent
  Distributed
Verification of Desired Properties (LTL)
Extending Distributed Model
  Rate Controlled Pacing
  Hysteresis Pacing
  Model the Noise, Diagnostics and ATR Mode
  Add More Parameters like Width and Amplitude

## Slide 29: Thank You
Questions ???
Contact – asankhaya@nus.edu.sg